## 1
|  |  | Incoming Inspection Report |  |  | Written by | Reviewed by | Approved by
|  |  | PRODUCT | Seegene STARlet |  | / / | / / | / /
Document Number |  |  |  |  | Serial Number
Incoming Inspection Instruction |  | SG-SI-STA-01, Rev.5 |  |  | Instrument | Serial No. | Seegene STARlet
Incoming Inspection Report |  | QFI-0803-ISS-01, Rev.6 |  |  | Computer | Serial No. | 5CG9171CF2
Inspection | Inspector |  |  |  |  | Result
| Start date |  | 2022-06-14
| Completed date |  | 2022-07-14
Hamilton S/W Version
Hamilton FV2 Version
Firmware Version | Pipette Ch (PX)
X-drive (XO)
Master (CO)
1. Appearance Inspection
1) Packaging Condition
Visual Inspection | Category |  | Factor | Determine to
| Shock Watch |  | Damage | ■ No □ Yes
| Binding |  | Installed | □ No ■ Yes
| Labels |  | Tagged | □ No ■ Yes
| Packaging Box |  | Damage | ■ No □ Yes
| Wooden Pallet |  | Damage | ■ No □ Yes
| Transport Jig |  | Installed | □ No ■ Yes
2) Appearance of instrument - Check if a scratch / damage / spots on below
Front
■ Pass
□ Fail
Top
■ Pass
□ Fail
Right
■ Pass
□ Fail
Left
■ Pass
□ Fail
Back
■ Pass
□ Fail

## 2
3) Instrument Accessories Kits
Category | Method | Factor | Result
Accessories | Visual Inspection | Damage | ■ No □ Yes
|  | Missing | ■ No □ Yes
2. Electrical Test
Category | Method | Acceptance Criteria | Result
Hi Potential | be based on manufacturer report | Insulation is normal Test potential: 1.5/2.3KV Time duration: 1 minute | ■ Pass □ Fail
Ground Resistance |  | ≤0.1 Ohm(Ω) | ■ Pass □ Fail
Power S/W Function | Visual Inspection | On/Off | ■ Pass □ Fail
Power LED On/Off | Visual Inspection | On/Off | ■ Pass □ Fail
3. Hardware Adjustment
Used Channel Calibration Tool Management No :
Used Barcode Carrier Management No :
Category | Method | Acceptance Criteria | Result
X-Arm | Adjust Arm Z using PIP | ≤+/-5 | Left : None Right : None
| Check Arm Diff | ≤+/-10
Pipette Channel | Adjust PIP Manual | Criteria | PX | dev // tilt | tilt | S/N
|  | X dev/tilt : ≤+/-20, Y tilt : ≤+/-10 | P1 | X : None // None | Y : None
|  | X dev/tilt : ≤+/-20, Y tilt : ≤+/-10 | P2 | X : None // None | Y : None
|  | X dev/tilt : ≤+/-20, Y tilt : ≤+/-10 | P3 | X : None // None | Y : None
|  | X dev/tilt : ≤+/-20, Y tilt : ≤+/-10 | P4 | X : None // None | Y : None
|  | X dev/tilt : ≤+/-20, Y tilt : ≤+/-10 | P5 | X : None // None | Y : None
|  | X dev/tilt : ≤+/-20, Y tilt : ≤+/-10 | P6 | X : None // None | Y : None
|  | X dev/tilt : ≤+/-20, Y tilt : ≤+/-10 | P7 | X : None // None | Y : None
|  | X dev/tilt : ≤+/-20, Y tilt : ≤+/-10 | P8 | X : None // None | Y : None
| Adjust PIP | □ Pass □ Fail
| Check PIP | X(1/100mm)≤+/-20, Y(1/100mm)≤+/-20 | P1 | X : 0 Y : 0
|  | Z(1/100mm)≤+/-30 |  | Z : 0
|  | X(1/100mm)≤+/-20, Y(1/100mm)≤+/-20 | P2 | X : 0 Y : 0
|  | Z(1/100mm)≤+/-30 |  | Z : 0
|  | X(1/100mm)≤+/-20, Y(1/100mm)≤+/-20 | P3 | X : 0 Y : 0
|  | Z(1/100mm)≤+/-30 |  | Z : 0
|  | X(1/100mm)≤+/-20, Y(1/100mm)≤+/-20 | P4 | X : 0 Y : 0
|  | Z(1/100mm)≤+/-30 |  | Z : 0
|  | X(1/100mm)≤+/-20, Y(1/100mm)≤+/-20 | P5 | X : 0 Y : 0
|  | Z(1/100mm)≤+/-30 |  | Z : 0
|  | X(1/100mm)≤+/-20, Y(1/100mm)≤+/-20 | P6 | X : 0 Y : 0
|  | Z(1/100mm)≤+/-30 |  | Z : 0
|  | X(1/100mm)≤+/-20, Y(1/100mm)≤+/-20 | P7 | X : 0 Y : 0
|  | Z(1/100mm)≤+/-30 |  | Z : 0
|  | X(1/100mm)≤+/-20, Y(1/100mm)≤+/-20 | P8 | X : 0 Y : 0
|  | Z(1/100mm)≤+/-30 |  | Z : 0
Autoload | Check Autoload | ■ Pass □ Fail
Noise | Unusual Sound generated | □ Pass ■ Fail
Remark

## 3
4. Performance Verification
Category |  | Acceptance Criteria |  |  |  |  | Measured Value
Cover Safety |  | Front door lock solenoid switch On/Off |  |  |  |  | ■ Pass □ Fail
Right side door lock solenoid switck On/Off
Barcode |  | Horizontal mode read Barcode |  |  |  |  | ■ Pass □ Fail
Vertical mode read Barcode
XYZ positioning |  | Status of Deviation in X Y Z direction(mm) ≤+/-1.5mm |  |  |  |  | ■ Pass □ Fail
HHS | HHS S/N | Temperature |  | Target 37℃, Tolerance ≤+/-1.0°C |  |  | ■ Pass □ Fail
Target 60℃, Tolerance ≤+/-1.5°C
Target 105℃, Tolerance ≤+/-2.5°C
|  | Shaking RPM |  | Target 200rpm/1000rpm/2400rpm, Tolerance +/-5% |  |  | ■ Pass □ Fail
Carrier Loading test on track |  | 2x, Part number 182085, TIP CARRIER, 5 TIP RACKS carrier |  |  |  |  | ■ Pass □ Fail
|  | 3x, Part number 173410, 3 Carrier for 32 Tubes carrier |  |  |  |  | ■ Pass □ Fail
|  | 1x, Part number 97319-01, ASSY, CAR, MFX, 6T, REAG, DW, MTP carrier |  |  |  |  | ■ Pass □ Fail
|  | 1x, Part number 97076-01. ASSY, CAR, 7T, HHS, 4MFX carrier |  |  |  |  | ■ Pass □ Fail
Volume Field Verification |  | P1 | Used Tip type : 300ul | Testing volume : 10ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  |  | Used Tip type : 1000ul | Testing volume : 1000ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  | P2 | Used Tip type : 300ul | Testing volume : 10ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  |  | Used Tip type : 1000ul | Testing volume : 1000ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  | P3 | Used Tip type : 300ul | Testing volume : 10ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  |  | Used Tip type : 1000ul | Testing volume : 1000ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  | P4 | Used Tip type : 300ul | Testing volume : 10ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  |  | Used Tip type : 1000ul | Testing volume : 1000ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  | P5 | Used Tip type : 300ul | Testing volume : 10ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  |  | Used Tip type : 1000ul | Testing volume : 1000ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  | P6 | Used Tip type : 300ul | Testing volume : 10ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  |  | Used Tip type : 1000ul | Testing volume : 1000ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  | P7 | Used Tip type : 300ul | Testing volume : 10ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  |  | Used Tip type : 1000ul | Testing volume : 1000ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  | P8 | Used Tip type : 300ul | Testing volume : 10ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
|  |  | Used Tip type : 1000ul | Testing volume : 1000ul | Accuracy | Tolerance ≤+/- 6.0% | ■ Pass □ Fail
[5 empty] | Precision | Tolerance ≤+/- 3.0% | ■ Pass □ Fail
5. Attachment
Attached document List | [6 empty] | Attached
Cover-Safety Verification Report | [6 empty] | ■
Barcode Verification Report | [6 empty] | ■
Position Verification Report | [6 empty] | ■
Declaration of Quality (Hamilton Heater Shaker ) | [6 empty] | ■
Declaration of Quality (Microlab STARlet IVD) | [6 empty] | ■
Measurement Protocol ML STAR IVD | [6 empty] | ■
STARLINE Electrical Safety Test Report | [6 empty] | ■
Remark

## 4
Seegene STARlet
Accessories Kits
Appendix > Accessories kits list |  |  |  | STARlet Serial number : | 5CG9171CF2
No | Part number | Description |  |  | Qty | Check
Accessories kit #1
1 | N/A | [Document] Unpacking Instructions, ML STAR |  |  | 1 | ■
2 | N/A | [Document] Installation Qualification, ML STAR, IVD |  |  | 1 | ■
3 | N/A | [Document] Final Configuration TEST STAR-Checklist |  |  | 1 | ■
4 | N/A | [Document] Measurement Protocol ML STAR IVD |  |  | 1 | ■
5 | N/A | [Document] EU Declaration of Conformity, ML STAR, IVD |  |  | 1 | ■
6 | N/A | [Document] Declaration of Quality, Replacement Autoload |  |  | 1 | ■
7 | N/A | [Document] Declaration of Quality, ML STAR, IVD |  |  | 1 | ■
8 | N/A | [Document] STARLINE Packing Checklist |  |  | 1 | ■
9 | N/A | [Document] Maintenance (Barcode, Position, Weekly) |  |  | 1 | ■
10 | N/A | [Document] Channel Adjust Value for Guldi Device |  |  | 1 | ■
11 | N/A | [Document] STARLINE Electrical Safety Test Report |  |  | 1 | ■
12 | 173240 | Loading Table, 30T |  |  | 1 | ■
13 | 185091-RNO | Side Cover Plexiglas right(Flip up) |  |  | 1 | ■
14 | 186571 | Side Cover Plexiglas left |  |  | 1 | ■
15 | 186572 | REAR Plexiglas Cover 30T(for STARlet) |  |  | 1 | ■
16 | 186573 | TOP Panel Plexiglas Cover 30T(for STARlet) |  |  | 1 | ■
17 | 190201 | LEFT REAR CORNER COVER (with Black) |  |  | 1 | ■
18 | 190202 | RIGHT REAR CORNER COVER (with Black) |  |  | 1 | ■
Accessories kit #2
19 | 96718-01 | ASSY, LIQ WASTE, 5L, 4L-SENSOR |  |  | 1 | ■
20 | 53873-07 | SOLID WASTE CONTAINER |  |  | 1 | ■
21 | 54698-01 | [USB]Liquid and Tip Waste System, User Guide, Rev B |  |  | 1 | ■
22 | 54802-01 | [CD] Solid-Liquid Waste Sensor V1.3, Rev B |  |  | 1 | ■
23 | 911265-01 | [USB] VENUS FOUR CD SW & MANUAL STAR IVD 4.5, Rev X1 |  |  | 1 | ■
24 | 182238 | 32 Inserts for 1.5ml Eppendorf Cups |  |  | 3 | ■
25 | 184089 | CORE Gripper (Screw x 4ea included) |  |  | 1 | ■
26 | 182136 | Teaching Needle(8ea) [LOT: Teaching_Needles ] |  |  | 1 | ■
27 | 53894-01 | SEPARATOR, MAGNETIC, NUCLEO |  |  | 1 | ■
28 | 58938-01 | Tip Eject Plate(DEFLECTOR, TIP, LIQUID WASTE) |  |  | 1 | ■
29 | 3892-01 | POWER CORD, STD, CEE 7/7 |  |  | 1 | ■
30 | 3892-05 | POWER CORD, USA, NEMA 5-15P |  |  | 1 | ■
31 | 355171 | CABLE USB 2.0 A-B 3M FERRIT |  |  | 1 | ■
32 | 363013 | FUSE 6.3AT |  |  | 2 | ■
33 | 420629 | SCREW M4x16 |  |  | 1 | ■
34 | 53866-01 | FITTING, QIK-CON, 3/8 IDX 3/8BARB |  |  | 1 | ■
35 | 54200-01 | ASSY, FILTER, AEROSOL, LIQ WASTE |  |  | 1 | ■
Transport materials
36 | N/A | (Channel support)X-Arm Holding brakets |  |  | 2 | ■
37 | 190148 | TRANSPORT SUPPORT BAR LEFT |  |  | 1 | ■
38 | 186675 | Support Bar on Top (for Top Plexiglas Cover) |  |  | 1 | ■
Accessories kit #3, Upper box
39 | 53686-01 | BAG, BIO-HAZARD, 14.5 x 8.2 x 18, 200/CS |  |  | 1 | ■
40 | 173410 | 3 Carrier for 32 Tubes(11x60 - 14x120mm) |  |  | 1 | ■
41 | 182085 | TIP CARRIER, 5 TIP RACKS |  |  | 2 | ■
42 | 97076-01 | ASSY, CAR, 7T, HHS, 4MFX |  |  | 1 | ■
|  | 97432-01 | PED, ADAPTER, SMALL TUBE, SEEG |  | 2 | ■
|  | 403453 | SCREW, FLHX M3X8, DIN7991, A2SS |  | 4 | ■
|  | 400263 | SCREW, SOHX, M3X8, DIN912, A2SS |  | 2 | ■
|  | N/A | HHS Base Plate |  | 1 | ■
43 | 97319-01 | ASSY, CAR, MFX, 6T, REAG, DW, MTP |  |  | 1 | ■
44 | 62964-03 | RENO KIT 3, STAR VERIFY KIT |  |  | 1 | ■
45 | 235969 | STD VOL CO-RE TIPS FIL 5x96 |  |  | 1 | ■
46 | 235970 | HIGH VOL EO-RE TIPS FIL 5x96 |  |  | 1 | ■
47 | 199034 | Hamiltion Heater & Shaker (HHS) [SN : ] |  |  | 1 | ■
48 | 99088-24 | [USB] Hamilton Heater Shaker Manual |  |  | 1 | ■
49 | N/A | Stop Barcode |  |  | 1 | ■
Remark